## heat_load
mechanism | Q [W] | Q [mW] | note
warm stage T [K] | 300
cold stage T [K] | 50
RADIATION (viewFactor) | 3.2555 | 3255.5228 | integral qr on cold walls
CONDUCTION (65-cable harness, CFD) | 1.0291 | 1029.0893 | coax k_eff, warm-cap wallHeatFlux
CONVECTION | 0 | 0 | vacuum, g=0
TOTAL | 4.2846 | 4284.6121 | radiation + conduction

## coax_conduction
coax | cable | n_cables | family | Q_CFD [W] | Q_target [W] | keff_fit_%
coax_L1 | SS_UT085 | 25 | drive | 0.3958 | 0.39579 | 0.127
coax_L2 | SS_UT085 | 25 | flux | 0.3958 | 0.39579 | 0.127
coax_L3 | SS_UT085 | 10 | readout+readin | 0.15832 | 0.15832 | 0.127
coax_L4 | SS_UT085 | 5 | pump | 0.079161 | 0.079158 | 0.127
TOTAL |  |  |  | 1.0291 | 1.029

## by_patch
region | patch | cold? | area [m2] | Q_rad [W] | Q_wall [W]
domain0 | warmPlate | no | 0.13745 | -1.9957 | 1.9958
domain0 | OVC | no | 0.86687 | -11.8144 | 11.8147
domain0 | shield_slave | yes | 0.58719 | 1.5105 | -1.5107
domain1 | warmPlate | no | 0.24327 | -3.7185 | 3.7186
domain1 | coldPlate | yes | 0.24313 | 0.78487 | -0.78487
domain1 | shield | yes | 0.56742 | 0.92168 | -0.92178
domain1 | domain1_to_coax_L2 | yes | 0.024319 | 0 | 3.614e-06
domain1 | domain1_to_coax_L4 | yes | 0.024317 | 0 | 3.6117e-06
domain1 | domain1_to_coax_L1 | yes | 0.024317 | 0 | 3.6132e-06
domain1 | domain1_to_coax_L3 | yes | 0.024316 | 0 | 3.6115e-06
domain1 | shield_bottom | yes | 0.019765 | 0.038518 | -0.038518

## energy_balance
region | sum qr [W] | sum whf [W] | net [W]
domain0 | -12.2997 | 12.2997 | -4.5151e-09
domain1 | -1.9734 | 1.9734 | 4.5659e-09

## gates
gate | result | detail
no_solver_errors | PASS | log=log.chtMultiRegionSimpleFoam
energy_cons_domain0 | PASS | net sum(whf+qr)=-4.52e-09 W (0.0%)
energy_cons_domain1 | PASS | net sum(whf+qr)=+4.57e-09 W (0.0%)
radiation_active | PASS | radiation on cold = 3255.523 mW
G3.1_conduction_CFD_vs_harness | PASS | CFD=1029.09 mW target=1029.05 mW err=0.0%
G3.4_keff_fit | PASS | max k_eff poly(order 5) fit err = 0.13%
G3.5_harness_accounting | PASS | 65 lines; W/cable=15.832 mW; total cond=1029.09 mW

## NOTES
Full CFD vacuum-radiation CHT (p=0.001 Pa, g=0).
Radiation: viewFactor (CFD). Conduction: 65-line Krinner harness via coax k_eff(T) (CFD, G3.1=0%). Convection: 0 (vacuum).
Coax = lumped stand-in; k_eff=(n_k/A_model) Sum A_i k_i(T), UT-085-SS-SS stainless, NIST k(T), order-5 poly.
A_model calibrated per geometry so CFD axial conduction = analytic harness target.
Cooling budgets: 50K/77K stage ~30 W; 4K stage 0.7-1.5 W (Krinner 2019 / Raicu 2025).