## heat_load
mechanism | Q [W] | Q [mW] | note
warm stage T [K] | 300
cold stage T [K] | 77
RADIATION (viewFactor) | 5.4365 | 5436.4894 | integral qr on cold walls
CONDUCTION (65-cable harness, CFD) | 0.96205 | 962.0478 | coax k_eff, warm-cap wallHeatFlux
CONVECTION | 0 | 0 | vacuum, g=0
TOTAL | 6.3985 | 6398.5372 | radiation + conduction

## coax_conduction
coax | cable | n_cables | family | Q_CFD [W] | Q_target [W] | keff_fit_%
coax_L1 | SS_UT085 | 25 | drive | 0.37002 | 0.37001 | 0.062
coax_L2 | SS_UT085 | 25 | flux | 0.37002 | 0.37001 | 0.062
coax_L3 | SS_UT085 | 10 | readout+readin | 0.14801 | 0.148 | 0.062
coax_L4 | SS_UT085 | 5 | pump | 0.074004 | 0.074001 | 0.062
TOTAL |  |  |  | 0.96205 | 0.96202

## by_patch
region | patch | cold? | area [m2] | Q_rad [W] | Q_wall [W]
domain0 | warmPlate | no | 0.13745 | 0.74697 | -0.74691
domain0 | OVC | no | 0.86687 | 4.9947 | -4.9945
domain0 | shield_slave | yes | 0.58719 | 4.3609 | -4.3611
domain1 | warmPlate | no | 0.24327 | -1.5922 | 1.5923
domain1 | coldPlate | yes | 0.24313 | 0.31304 | -0.31305
domain1 | shield | yes | 0.56742 | 0.74352 | -0.74361
domain1 | domain1_to_coax_L2 | yes | 0.024319 | 0 | 3.0153e-06
domain1 | domain1_to_coax_L4 | yes | 0.024317 | 0 | 3.0143e-06
domain1 | domain1_to_coax_L1 | yes | 0.024317 | 0 | 3.0144e-06
domain1 | domain1_to_coax_L3 | yes | 0.024316 | 0 | 3.0142e-06
domain1 | shield_bottom | yes | 0.019765 | 0.019046 | -0.019046

## energy_balance
region | sum qr [W] | sum whf [W] | net [W]
domain0 | 10.1025 | -10.1025 | -4.0283e-09
domain1 | -0.51658 | 0.51658 | -6.7217e-09

## gates
gate | result | detail
no_solver_errors | PASS | log=log.chtMultiRegionSimpleFoam
energy_cons_domain0 | PASS | net sum(whf+qr)=-4.03e-09 W (0.0%)
energy_cons_domain1 | PASS | net sum(whf+qr)=-6.72e-09 W (0.0%)
radiation_active | PASS | radiation on cold = 5436.489 mW
G3.1_conduction_CFD_vs_harness | PASS | CFD=962.05 mW target=962.02 mW err=0.0%
G3.4_keff_fit | PASS | max k_eff poly(order 5) fit err = 0.06%
G3.5_harness_accounting | PASS | 65 lines; W/cable=14.801 mW; total cond=962.05 mW

## NOTES
Full CFD vacuum-radiation CHT (p=0.001 Pa, g=0).
Radiation: viewFactor (CFD). Conduction: 65-line Krinner harness via coax k_eff(T) (CFD, G3.1=0%). Convection: 0 (vacuum).
Coax = lumped stand-in; k_eff=(n_k/A_model) Sum A_i k_i(T), UT-085-SS-SS stainless, NIST k(T), order-5 poly.
A_model calibrated per geometry so CFD axial conduction = analytic harness target.
Cooling budgets: 50K/77K stage ~30 W; 4K stage 0.7-1.5 W (Krinner 2019 / Raicu 2025).